## mPIMA HIV VL
ASLM is not responsible for the integrity of the data included herein after data files have been exported from the Diagnostic Evidence Hub
Author(s) | Publication year | Publication type | Type of study | Country | Study period | Intervention (assay) | Comparator (assay) | Sample Type | Sample size (number of specimens analyzed) | Study setting | Patient/study population | Where sample was collected | Where sample was tested | Device users/operators, type and number (if available) | Concordance, agreement, or correlation | Bland Altman's mean of difference (log copies per ml) | Sensitivity | Specificity | Error/failure rate | Citation
Bwana P, et al | 2019 | Journal article | Performance evaluation and usability study | Kenya | June 2018 - Nov 2018 | mPIMA HIV-1/2 Viral Load | Abbott M200 HIV-1 Real Time | Plasma | 566 | Alupe, Nambale, and Matayos Sub Country hospitals and Siaya County Referral Hospital | HIV-infected patients ≥ 13 years | Alupe, Nambale, and Matayos Sub Country hospitals and Siaya County Referral Hospital | Alupe, Nambale, and Matayos Sub Country hospitals and Siaya County Referral Hospital | 7 lab technologists | Agreement: 95.86% at 1000 copies/ml threshhold 97.77% at 800 copies/ml 72.61% at detection limit for each device R2 = 0.8538 | 0.16 (95% LOA: - 0.514; 0.835) | nr/nd | nr/nd | 7.77% | Bwana P, Ageng'o J, Danda J, Mbugua J, Handa A, Mwau M. Performance and usability of mPIMA™ HIV 1/2 viral load test in point of care settings in Kenya. J Clin Virol. 2019;121:104202. PMID: 31715524; DOI: 10.1016/j.jcv.2019.104202.
Mariani D, et al | 2020 | Journal article | Performance evaluation | Brazil | nr/nd | mPIMA HIV-1/2 Viral Load | Abbott M200 HIV-1 Real Time | Plasma | 413 | Hospital Universitario Graffee Guinle (Rio de Janeiro, Brazil) | HIV-infected patients failing ART (age not specified) | Hospital Universitario Graffee Guinle (Rio de Janeiro, Brazil) | nr/nd | nr/nd | Concordance: 88.9 % (95 % CI 85.4-91.7%) at 1000 copies/ml threshold Pearson r = 0.925 | - 0.26 (95% LOA: - 0.92; 0.40) | 97.1% (95% CI: 94.2-98.8%) | 76.9% (95% CI: 69.8 - 83.1%) | nr/nd | Mariani D, de Azevedo MCVM, Vasconcellos I, et al. The performance of a new point-of-care HIV virus load technology to identify patients failing antiretroviral treatment. J Clin Virol. 2020;122:104212. PMID: 31765957; DOI: 10.1016/j.jcv.2019.104212
nr/nd = not reported/no data